Performance of mRNA and DNA Based High-Risk Human Papillomavirus Assays in Detection of High-Grade Cervical Lesions

Running headline: Evaluation of hrHPV assays

Elina Virtanen (MSc)1, Ilkka Kalliala (MD, PhD)2,3, Tadeusz Dyba (ScD)4, Pekka Nieminen (MD, PhD)2, Eeva Auvinen (PhD)1

1 Virology, University of Helsinki and Helsinki University Hospital, Helsinki, Finland

2 Gynecology and Obstetrics, University of Helsinki and Helsinki University Hospital, Helsinki, Finland

3 Institute of Reproduction and Developmental Biology, Department of Surgery & Cancer, Imperial College London, United Kingdom

4 Public Health – Cancer Policy Support, Institute for Health and Consumer Protection, DG Joint Research Centre, European Commission, Ispra, Italy

Corresponding author details: 	Elina Virtanen

Medicum, Department of Virology

POB 21 (Haartmaninkatu 3), 00014 Helsinki

+358 40 838 4010

elina.i.virtanen@helsinki.fi

Conflict of interest

None of the authors have conflicts of interests to declare.

Abstract

Introduction. The aim was to assess the performance of two commercial assays for the detection of high-risk human papillomavirus (hrHPV): Aptima HPV Assay (Hologic, Inc., Marlborough, MA, USA) detecting mRNA of 14 different hrHPV types, and Hybrid Capture 2 HPV DNA test (HC2; Qiagen, Gaithersburg, MD, USA) that detects the DNA of 13 different hrHPV types. Test performance was compared in the settings of a standard colposcopy clinic, among the regular patient flow.

Material and Methods. Two separate cervical cell samples for Aptima and HC2 testing were collected from women referred to colposcopy or cervical follow-up visit. Altogether 481 paired samples were analyzed and all positive samples were also tested using the Aptima HPV 16 18/45 Genotype Assay. Results from the two assays were compared directly and with stratification by histology and cytology from the same sampling visit.

Results. The overall agreement between HC2 and Aptima assays was 92.9% (Kappa coefficient of 0.855). The sensitivity and specificity of the assays in detecting CIN2+ were 92.5% and 58.2% for HC2, and 94.0% and 59.3% for Aptima, respectively. No significant differences between the assays were found (p-values >0.5). Both assays detected all CIN3 (n = 30) and carcinoma (n = 2) cases.

Conclusions. The mRNA-based Aptima assay and the extensively studied DNA-based HC2 test performed equally well in detecting high-grade cervical lesions. Our data contribute to the growing evidence base indicating that the mRNA-based Aptima assay could be used for the triage of patients with HPV-associated cervical disease.

Keywords

Cervical intraepithelial neoplasia, CIN, diagnostics, HPV, human papillomavirus, test performance

Abbreviations

ASC-US		atypical squamous cells of undetermined significance

CIN			cervical intraepithelial neoplasia

hrHPV		high-risk human papillomavirus

LSIL			low-grade squamous intraepithelial lesion

Key Message

The study showed that the performance of mRNA-based Aptima assay did not differ from DNA-based HC2 test in the detection of cervical HPV associated lesions. The evaluated hrHPV assays were found equal in performance and more sensitive than conventional cytology.

Introduction

Human papillomaviruses (HPV) are double-stranded DNA viruses infecting epithelial cells of skin and mucosa (1). Mucosal HPV genotypes are divided into low-risk (lrHPV) and high-risk (hrHPV) groups based on their association with severe disease or cancer (1). HPV is best known for its association with cervical cancer and high-grade cervical lesions (CIN2+), which are caused by a persistent hrHPV infection (2,3). In cervical cancer screening programs the cytology-based Papanicolaou (Pap) test has been traditionally used. Recently it has been shown that also hrHPV testing can be used in primary cervical cancer screening of 30-year-old and older women, and also as a triage test of atypical squamous cells of undetermined significance (ASC-US) cytology, as well as test of cure after treatment of cervical intraepithelial neoplasia (CIN) (4,5). HrHPV testing has also proven more sensitive than the Pap test (6-8) and has a higher negative predictive value, which means that a negative HPV test result indicates a very low risk of developing cervical cancer within the next 5–10 years (9-12). The challenge with HPV testing is the lower specificity as compared to cytology in age group below 30 years of age (6).

In this study we compared the performance of two commercial hrHPV assays in the detection of high-grade cervical lesions and cancer: Hybrid Capture 2 (HC2; Qiagen, Gaithersburg, MD, USA) and Aptima HPV Assay (Hologic, Inc., Marlborough, MA, USA). Test performance was compared in a study population of a colposcopy clinic, comprising women whose cytological and histological findings vary from normal to LSIL or CIN3 respectively. No specific selection criteria were applied; instead we wanted to evaluate the tests in the practical settings of a standard colposcopy clinic, and among the regular patient flow where the use of hrHPV testing could be considered.

Material and Methods

Altogether 481 paired cervical cell samples were collected from Finnish women who visited the colposcopy clinic at the Department of Gynecology at the Helsinki University Jorvi Hospital between September 2013 and March 2014. The cohort comprised women with recent cervical findings as well as those who had already been treated and monitored for a longer period of time – altogether representing a typical population attending a tertiary-care colposcopy clinic. Two separate cell samples were collected, a few seconds apart, using Digene Cervical Sampler and Specimen Transport device (Qiagen) and Aptima Cervical Specimen Collection and Transport device (Hologic). The samples were taken in random order, and during the same visit also cytological and/or histological samples were collected based on clinical decision, independent of study sample collection. Conventional Papanicolaou test was used for cytology. All cytological and histological analyses were carried out at the Department of Pathology of Helsinki University Hospital Laboratory. The Ethics Committee of the Hospital District of Helsinki and Uusimaa approved the collection of samples for this study on 29th of April in 2013 (reference number 130/13/03/03/2013) and written informed consent was obtained from all patients.

Cervical samples were tested using two different methods: the manual Hybrid Capture 2 (HC2) HPV DNA test (Qiagen) that detects the DNA of 13 different hrHPV types (16, 18, 31, 33, 35, 39, 45, 51, 52, 56, 58, 59, 68) and the Aptima HPV Assay (Hologic) with the fully automated Panther system (Hologic) detecting the mRNA of 14 different hrHPV types (those mentioned above and additionally HPV 66). Both assays are FDA-approved and CE-IVD-marked. The assays were performed according to the manufacturer’s instructions. All samples found positive with either assay were also genotyped using the Aptima HPV 16 18/45 Genotype Assay (Hologic). To further verify test performance in our hands, a subset of HPV negative samples was examined with the genotyping assay as well. All HPV analyses were carried out at the Department of Virology and Immunology of Helsinki University Hospital Laboratory. In Aptima assay, the signal to cut-off (S/CO) value of ≥ 0.5 was interpreted as positive. In HC2 assay, the limit of positivity was the relative light unit per cut-off (RLU/CO) value ≥ 1. All HC2 samples with initial RLU/CO values between 0.8 and 1.5 were retested once or twice to confirm the result, according to the manufacturer’s recommendation. The performance of the two hrHPV assays as well as cytology in the detection of histologically verifiable diseases was assessed. Samples for HPV testing, cytology and histology were collected at the same visit.

Statistical analysis

Data analysis was performed using CIN2 or more severe lesion (CIN2+) as a cut-off for presence of the disease. The results were also stratified according to patient age, reason for colposcopy and cytological status. For HC2, Aptima and cytology sensitivity, specificity, predictive values and their 95.0% confidence intervals (CI) were calculated assuming binomial distribution of the test’s results. The p-values for comparing sensitivity and specificity of tests were calculated using McNemar's test (13). The agreement between the two hrHPV assays was evaluated by Kappa statistics (14). All statistical analyses were performed with STATA software (StataCorp 2013, Stata Statistical Software: release 13.1, College Station, TX, USA).

Results

The age of the patients ranged from 18 to 86 years. The majority of women (59.4%) were 35–65 years of age, whereas 32.7% were 18–34 years and 7.9% were 66–86 years of age. Of all HPV tested women 196 (40.8%) attended HPV sampling because of an abnormal Pap test, and from 197 women (41.0%) a biopsy had been taken on the preceding visit before HPV sampling due to earlier suspicion of HPV related disease.  For the remaining 87 women (18.1%) the reason for HPV sampling was continuing surveillance because of previous findings, other specified reason (for example pain in lower abdomen or bleeding), or the reason was not clearly specified in the patient registry. The time between the preceding visit and HPV sampling ranged from less than one month to one year.

Altogether 481 paired samples were analyzed using both Aptima and HC2 methods. One Aptima sample was insufficient, and that sample pair was thus excluded from the analysis leaving 480 paired samples for the analysis. In total, 202 samples (42.1%) tested positive using the HC2 method and 206 (42.9%) by the Aptima method (Table 1). There were 34 (7.1%) discrepant samples of which 15 were HC2 positive but Aptima negative, and 19 Aptima positive but HC2 negative (Table 1). The overall agreement between the assays was 92.9% (Kappa coefficient of 0.855).

HrHPV assay results as well as sensitivities, specificities, and positive and negative predictive values within histological groups are presented in Tables 2 and 3. The sensitivity and specificity to detect CIN2+ was 92.5% and 58.2% for HC2, and 94.0% and 59.3% for Aptima, respectively. No significant differences in test performance between the two hrHPV assays were found (p-values >0.5). Specifically, five CIN2 cases (14.3%) remained hrHPV negative by the HC2 assay and four (11.4%) by the Aptima assay. Importantly, all CIN3 cases (n = 30) and the two carcinomas found were detected by both hrHPV assays. Of the samples giving discrepant results with the two hrHPV assays six had normal histology, 15 had CIN1 grade lesions, three had CIN2 grade lesions and 10 had no histological reference. Of the three discrepant CIN2 cases one was positive only with HC2 and two with Aptima.

Results were also analyzed separately according to the reason for colposcopy referral (Table 4). The study population was divided into three groups based on the preceding visit before HPV sampling. These groups were: 1) patients who had Pap test suggestive of cervical lesion, 2) patients who had a treatment or control visit at which histology was obtained and 3) other patients (long time follow-up with normal findings, specified reasons such as bleeding or pain in lower abdomen, the preceding visit was more than one year before HPV sampling, or no data was available about the previous visit). In every group the predictive values were the same for both assays or slightly higher for Aptima. Aptima was also 2.1 percentage points more specific when the reason for colposcopy was previous Pap test suggestive of cervical lesion, and 8.4 percentage points more sensitive for women whose histological status had been confirmed on the preceding visit. However, these differences were not statistically significant (p-values >0.5). In the group of other reasons for colposcopy the performance was identical for both assays and both sensitivity and negative predictive value reached 100.0%.

HrHPV results with cytological findings at the HPV sampling visit are shown in Table 5. The rates of positive and negative hrHPV findings in the various cytological groups were similar for both hrHPV assays. The positivity rate increased from 20.7% to 95.0% from normal towards ASC-H and more severe findings. The performance of hrHPV assays was similar also when compared to cytological findings at the HPV sampling visit (Table 6). Among normal and non-severe (ASC-US and LSIL) findings, both the sensitivity and the specificity were slightly higher for Aptima but the differences were not statistically significant (p-values >0.5). In ASC-H and more severe findings the sensitivity was 100.0% and the specificity 15.8% for both assays.

Sensitivity, specificity and positive and negative predictive values with 95.0% CI for the Pap test in the detection of CIN2+ are shown in Supporting Information, table S1. The sensitivity and specificity of cytology in detecting CIN2+ was 86.2% and 52.0% when using ASC-US as a cut-off, and 75.4% and 72.7% when LSIL was used as a cut-off. HrHPV testing was found more sensitive (p-value 0.003 for Aptima and 0.002 for HC2) in detecting CIN2+ than cytology when LSIL was used as a cut-off. However, cytology was more specific (p-value <0.001 for both assays), but at the cost of decreased sensitivity. Although both hrHPV assays demonstrated slightly higher sensitivity and specificity in detecting CIN2+ than cytology when ASC-US was used as a cut-off, only the better specificity of the Aptima assay was statistically significant (p-values for the sensitivity: 0.132 for Aptima and 0.206 for HC2, for the specificity: 0.033 for Aptima and 0.077 for HC2).

Test parameters were also calculated separately for two age groups, <30 and ≥30 years (Supporting Information, tables S1 and S2). The sensitivities for hrHPV assays were very similar between the two age groups but slightly higher in the younger age group (Table S2). For the Aptima assay the sensitivity reached 100.0% in women younger than 30. As for Pap test the sensitivities were higher in the older age group and the differences were more evident (Table S1). For both hrHPV and Pap tests there were clear differences in specificities between the age groups and for each testing method the specificities were higher in women aged 30 or older. The negative predictive values (NPV) for detecting CIN2+ were higher for hrHPV assays than for the Pap test in all age groups. The NPVs for the Pap test using either ASC-US or LSIL cut-off were more than 12 percentage points lower in women below 30 years than in the older age group. For hrHPV assays there were only minor differences between age groups and the NPV of the Aptima assay reached 100.0% in the younger age group. The positive predictive values (PPV) in detecting CIN2+ remained fairly low for both hrHPV and Pap tests in all age groups.

Aptima genotyping assay was performed for all samples found positive using either hrHPV assay. Of all samples in this study 10.2% were positive for HPV16 and only 4.6% for HPV18/45 (Supporting Information, table S3). One sample which was found positive in both Aptima and HC2 assays tested positive for both HPV16 and HPV18/45. The samples that were initially Aptima positive but HC2 negative (n = 19) included one sample which was positive for HPV18/45, two which were positive for HPV16, and 16 samples which were positive for other hrHPV genotypes. All initially Aptima negative but HC2 positive samples (n = 15) and the tested subset of samples that were initially negative with both methods (n = 55) remained negative also in the genotyping assay.

HrHPV was found in 88.6% of CIN2 grade lesions (Supporting Information, table S3). The majority (48.6%) were hrHPV types other than 16 or 18/45, whereas HPV16 was found in 31.4% and HPV18/45 in 8.6% of CIN2 lesions, respectively. High-risk HPV was found in all CIN3 lesions: HPV16 was detected in 40.0%, HPV18/45 in 13.3%, and other HrHPV types in 46.7% of CIN3 lesions. Of the two carcinomas one was positive for HPV16 and the other for HPV18/45. HrHPV types were also found in 31.2% of samples with normal histology. The majority of these (70.8%) were other hrHPV types than 16 or 18/45.

Discussion

The purpose of this study was to assess the performance of HC2 and Aptima hrHPV assays in the detection of high-grade cervical epithelial lesions among the routine patient flow at a tertiary-care colposcopy clinic. Cytological samples were assessed as well, and all results were compared to histological findings from biopsies collected at the same visit. Similar sensitivity, specificity, PPV and NPV were shown for both hrHPV assays. However, significant improvement in sensitivity was obtained as compared to the conventional Pap test. Because both hrHPV assays were found to be of equal performance, individual requirements of the laboratory may thus be decisive in choosing the appropriate hrHPV assay.

The overall agreement between the hrHPV assays was very high (92.9%) and similar rates have been reported in other studies, where the overall agreement has been between 88–97% (15-17). Overall both hrHPV assays performed with higher sensitivity and higher NPV than conventional cytology. Improvement of assay sensitivity often brings about increasing rate of false positive findings, decreasing test specificity. This is particularly evident in hrHPV testing, where test positivity is common in the absence of disease, especially in younger age groups. Many HPV infections among young women are transient, and the proportion of persistent infections increases in older age groups (18). In the case of mild disease spontaneous clearance often takes place (19,20). As seen in our study HrHPV testing is highly sensitive among young women but the specificity is clearly lower as compared to the older age group. Test specificity increases among women more than 30 years of age whose HPV infections are more likely to be persistent and cause disease. Still, among younger women the very high sensitivity and NPV add value to the negative test result indicating a very low risk of developing cervical disease and allowing longer screening intervals. However, because of frequent transient HPV infections among younger women, HPV testing is generally recommended for primary cervical cancer screening only among 30-year-old and older women (21).

Screening for cervical cancer using hrHPV testing instead of Pap aims at risk recognition, i.e. finding true or potential disease. The high NPV of hrHPV testing allows longer screening intervals with similar or lower risk of developing CIN3+ lesions as compared to Pap test (9,11,22). The lower NPV of cytology among women below 30 years of age shows that a negative Pap result does not necessarily indicate the absence of the disease. Based on different studies screening intervals of as much as 5–10 years could be considered in programs based on hrHPV screening (9-12,22,23). Effectiveness using primary HPV screening with cytology triage and five-year screening interval has been convincingly shown in the Finnish national screening program (22).

Mild lesions are often associated with low-risk HPV types, which cause a very low or non-existing risk for CIN3+ lesions and cervical cancer (24). The HPV assays evaluated in this study detect only high-risk HPV types. Thus lesions caused by low-risk types would have remained negative, although HC2 is known to cross-react with lrHPV types (25). The cross-reactivity of the Aptima assay has not yet been assessed in large clinical materials. Altogether five CIN2 cases remained negative using the HC2 test and four using the Aptima assay. Some CIN2 lesions are caused by low-risk genotypes with low risk of progression into cervical cancer (21,24,26,27), which may partially explain this finding. Overall, the sensitivity of hrHPV testing for CIN2+ was higher than that of Pap, and thus fewer severe cases will be missed by hrHPV than Pap testing. Importantly, all CIN3+ cases were found positive with both hrHPV assays in all age groups.

Consensus prevails on HPV16 being the genotype of highest risk, followed by HPV18 and HPV45 and possibly some other high-risk types. This was also seen in our study material where the frequency of HPV16 and HPV18/45 positivity clearly increased towards more severe disease, whereas other hrHPV types were found in all histological groups except for the two carcinomas. The overall prevalence of HPV16 (10.2%) and HPV18/45 (4.6%) was fairly low in our study population but in agreement with the previously published genotype distribution in the Finnish screening population showing a lower prevalence of genotypes 16, 18 and 45 in particular (28). The Aptima Genotype Assay provides a triaging option for HPV positive women in the future but as yet the European guidelines for cervical cancer screening do not specifically recommend the inclusion of HPV genotyping in patient management (29).

Previous studies have shown better specificity for Aptima than for HC2 (15-17,30-32). In both triage of ASC-US and LSIL, Aptima is as sensitive but more specific than HC2 for detecting cervical pre-cancer. Only Aptima has been considered useful in triage of LSIL (5). In the present study we were not able to reproduce this difference, instead the specificity was very similar for both hrHPV methods. As we found the HC2 and Aptima assays perform equally in detecting high-grade cervical lesions, the practical test properties may become decisive when choosing a HPV test for clinical use. These include the hands-on time required, additional testing opportunities including genotyping, as well as the required throughput and instrumentation available in the laboratory. Automation is also available for HC2 (Rapid Capture System and Hybrid Capture 2 Modular System, Qiagen), which diminishes the amount of manual work and shortens turn-around time. The benefits of Aptima with Panther system include flexible throughput, reduced hands-on time and full automation. The inclusion of an automated and highly sensitive hrHPV testing method, such as Aptima, could be used to increase the efficiency of cervical cancer screening and further reduce cervical cancer incidence.

The guidelines and criteria for HPV testing in primary cervical cancer screening have taken into account only DNA-based tests so far (21,33), but the Aptima HPV assay also fulfills the cross-sectional clinical and reproducibility criteria of international guidelines for human papillomavirus test requirements for cervical cancer screening (34). Our study shows that the performance of mRNA-based Aptima assay did not differ from the widely used and well-studied DNA-based HC2 test. Our data contribute to the growing evidence base on the Aptima HPV mRNA assay and indicate that it could be used for the triage of patients with the indication of HPV-associated cervical disease.

Acknowledgments

We thank the staff of the chlamydia laboratory at the Department of Virology and Immunology of the Helsinki University Hospital Laboratory for performing the Aptima analyses.

We also want to thank Hologic for providing a major proportion of the Aptima assays used in this study.

Funding statement

This study was funded by a research and development grant of the Helsinki University Hospital Laboratory (Eeva Auvinen).

References

1.	zur Hausen H. Papillomaviruses and cancer: from basic studies to clinical application. Nat Rev Cancer. 2002;2(5):342-350.

2.	Munoz N, Bosch FX, de Sanjose S, Herrero R, Castellsague X, Shah KV, et al. Epidemiologic classification of human papillomavirus types associated with cervical cancer. N Engl J Med. 2003;348(6):518-527.

3.	Walboomers JM, Jacobs MV, Manos MM, Bosch FX, Kummer JA, Shah KV, et al. Human papillomavirus is a necessary cause of invasive cervical cancer worldwide. J Pathol. 1999;189(1):12-19.

4.	Arbyn M, Ronco G, Anttila A, Meijer CJ, Poljak M, Ogilvie G, et al. Evidence regarding human papillomavirus testing in secondary prevention of cervical cancer. Vaccine. 2012;30 Suppl 5:F88-99.

5.	Arbyn M, Roelens J, Cuschieri K, Cuzick J, Szarewski A, Ratnam S, et al. The APTIMA HPV assay versus the Hybrid Capture 2 test in triage of women with ASC-US or LSIL cervical cytology: a meta-analysis of the diagnostic accuracy. Int J Cancer. 2013;132(1):101-108.

6.	Cuzick J, Clavel C, Petry KU, Meijer CJ, Hoyer H, Ratnam S, et al. Overview of the European and North American studies on HPV testing in primary cervical cancer screening. Int J Cancer. 2006;119(5):1095-1101.

7.	Koliopoulos G, Arbyn M, Martin-Hirsch P, Kyrgiou M, Prendiville W, Paraskevaidis E. Diagnostic accuracy of human papillomavirus testing in primary cervical screening: a systematic review and meta-analysis of non-randomized studies. Gynecol Oncol. 2007;104(1):232-246.

8.	Ronco G, Dillner J, Elfstrom KM, Tunesi S, Snijders PJ, Arbyn M, et al. Efficacy of HPV-based screening for prevention of invasive cervical cancer: follow-up of four European randomised controlled trials. Lancet. 2014;383(9916):524-532.

9.	Dillner J, Rebolj M, Birembaut P, Petry KU, Szarewski A, Munk C, et al. Long term predictive values of cytology and human papillomavirus testing in cervical cancer screening: joint European cohort study. BMJ. 2008;337:a1754.

10.	Elfstrom KM, Smelov V, Johansson AL, Eklund C, Naucler P, Arnheim-Dahlstrom L, et al. Long term duration of protective effect for HPV negative women: follow-up of primary HPV screening randomised controlled trial. BMJ. 2014;348:g130.

11.	Khan MJ, Castle PE, Lorincz AT, Wacholder S, Sherman M, Scott DR, et al. The elevated 10-year risk of cervical precancer and cancer in women with human papillomavirus (HPV) type 16 or 18 and the possible utility of type-specific HPV testing in clinical practice. J Natl Cancer Inst. 2005;97(14):1072-1079.

12.	Vink MA, Bogaards JA, Meijer CJ, Berkhof J. Primary human papillomavirus DNA screening for cervical cancer prevention: Can the screening interval be safely extended? Int J Cancer. 2015;137(2):420-427.

13.	McNemar Q. Note on the sampling error of the difference between correlated proportions or percentages. Psychometrika. 1947;12(2):153-157.

14.	Cohen J. A Coefficient of Agreement for Nominal Scales. Educ Psychol Meas. 1960;20(1):37-46.

15.	Clad A, Reuschenbach M, Weinschenk J, Grote R, Rahmsdorf J, Freudenberg N. Performance of the Aptima high-risk human papillomavirus mRNA assay in a referral population in comparison with Hybrid Capture 2 and cytology. J Clin Microbiol. 2011;49(3):1071-1076.

16.	Iftner T, Becker S, Neis KJ, Castanon A, Iftner A, Holz B, et al. Head-to-Head Comparison of the RNA-Based Aptima Human Papillomavirus (HPV) Assay and the DNA-Based Hybrid Capture 2 HPV Test in a Routine Screening Population of Women Aged 30 to 60 Years in Germany. J Clin Microbiol. 2015;53(8):2509-2516.

17.	Reuschenbach M, Clad A, von Knebel Doeberitz C, Wentzensen N, Rahmsdorf J, Schaffrath F, et al. Performance of p16INK4a-cytology, HPV mRNA, and HPV DNA testing to identify high grade cervical dysplasia in women with abnormal screening results. Gynecol Oncol. 2010;119(1):98-105.

18.	Rodriguez AC, Schiffman M, Herrero R, Hildesheim A, Bratti C, Sherman ME, et al. Longitudinal study of human papillomavirus persistence and cervical intraepithelial neoplasia grade 2/3: critical role of duration of infection. J Natl Cancer Inst. 2010;102(5):315-324.

19.	Ho GY, Bierman R, Beardsley L, Chang CJ, Burk RD. Natural history of cervicovaginal papillomavirus infection in young women. N Engl J Med. 1998;338(7):423-428.

20.	Moscicki AB, Shiboski S, Hills NK, Powell KJ, Jay N, Hanson EN, et al. Regression of low-grade squamous intra-epithelial lesions in young women. Lancet. 2004;364(9446):1678-1683.

21.	Meijer CJ, Berkhof J, Castle PE, Hesselink AT, Franco EL, Ronco G, et al. Guidelines for human papillomavirus DNA test requirements for primary cervical cancer screening in women 30 years and older. Int J Cancer. 2009;124(3):516-520.

22.	Anttila A, Kotaniemi-Talonen L, Leinonen M, Hakama M, Laurila P, Tarkkanen J, et al. Rate of cervical cancer, severe intraepithelial neoplasia, and adenocarcinoma in situ in primary HPV DNA screening with cytology triage: randomised study within organised screening programme. BMJ. 2010;340:c1804.

23.	Schiffman M, Glass AG, Wentzensen N, Rush BB, Castle PE, Scott DR, et al. A long-term prospective study of type-specific human papillomavirus infection and risk of cervical neoplasia among 20,000 women in the Portland Kaiser Cohort Study. Cancer Epidemiol Biomarkers Prev. 2011;20(7):1398-1409.

24.	Snijders PJ, Steenbergen RD, Heideman DA, Meijer CJ. HPV-mediated cervical carcinogenesis: concepts and clinical implications. J Pathol. 2006;208(2):152-164.

25.	Castle PE, Schiffman M, Burk RD, Wacholder S, Hildesheim A, Herrero R, et al. Restricted cross-reactivity of hybrid capture 2 with nononcogenic human papillomavirus types. Cancer Epidemiol Biomarkers Prev. 2002;11(11):1394-1399.

26.	Petry KU, Cox JT, Johnson K, Quint W, Ridder R, Sideri M, et al. Evaluating HPV-negative CIN2+ in the ATHENA trial. Int J Cancer. 2016;138(12):2932-2939.

27.	Arbyn M, Tommasino M, Depuydt C, Dillner J. Are 20 human papillomavirus types causing cervical cancer? J Pathol. 2014;234(4):431-435.

28.	Leinonen MK, Anttila A, Malila N, Dillner J, Forslund O, Nieminen P. Type- and age-specific distribution of human papillomavirus in women attending cervical cancer screening in Finland. Br J Cancer. 2013;109(11):2941-2950.

29.	Anttila A, Arbyn M, De Vuyst H, Dillner J, Dillner L, Franceschi S, et al. European guidelines for quality assurance in cervical cancer screening. Second Edition Supplements ed. Luxembourg: Publications Office of the European Union, 2015.

30.	Dockter J, Schroder A, Hill C, Guzenski L, Monsonego J, Giachetti C. Clinical performance of the APTIMA HPV Assay for the detection of high-risk HPV and high-grade cervical lesions. J Clin Virol. 2009;45 Suppl 1:S55-61.

31.	Rebolj M, Bonde J, Preisler S, Ejegod D, Rygaard C, Lynge E. Human Papillomavirus Assays and Cytology in Primary Cervical Screening of Women Aged 30 Years and Above. PLoS One. 2016;11(1):e0147326.

32.	Szarewski A, Mesher D, Cadman L, Austin J, Ashdown-Barr L, Ho L, et al. Comparison of seven tests for high-grade cervical intraepithelial neoplasia in women with abnormal smears: the Predictors 2 study. J Clin Microbiol. 2012;50(6):1867-1873.

33.	Arbyn M, Snijders PJ, Meijer CJ, Berkhof J, Cuschieri K, Kocjan BJ, et al. Which high-risk HPV assays fulfil criteria for use in primary cervical cancer screening? Clin Microbiol Infect. 2015;21(9):817-826.

34.	Heideman DA, Hesselink AT, van Kemenade FJ, Iftner T, Berkhof J, Topal F, et al. The Aptima HPV assay fulfills the cross-sectional clinical and reproducibility criteria of international guidelines for human papillomavirus test requirements for cervical screening. J Clin Microbiol. 2013;51(11):3653-3657.

Supporting Information legends:

Table S1. Sensitivity, specificity and positive and negative predictive values (95.0% CI) of Pap test in the detection of CIN2+ in age groups below and above 30 years of age (N = 330 samples with histology) for two different thresholds of test positivity (ASC-US+ and LSIL+).

Table S2. Sensitivity, specificity and positive and negative predictive values (95.0% CI) of the hrHPV tests in the detection of CIN2+ in age groups below and above 30 years of age (N = 330 samples with histology).

Table S3. Genotyping results and histological findings of Aptima hrHPV positive samples.

Legends of tables:

Table 1. HrHPV test results.

Table 2. HC2 results and histological status of the patients at the HPV sampling visit (N = 330 samples with histology). Performance parameters are shown for CIN2+.

Table 3. Aptima results and histological status of the patients at the HPV sampling visit (N = 330 samples with histology). Performance parameters are shown for CIN2+.

Table 4. Sensitivity, specificity and positive and negative predictive values of the hrHPV assays for the detection of CIN2+ stratified according to the reason for colposcopy.

Table 5. HrHPV results and cytological status of the patients at the HPV sampling visit (N = 480).

Table 6. Sensitivity, specificity and positive and negative predictive values of the hrHPV assays for the detection of CIN2+ in different cytological groups.

Tables:

Table 1. HrHPV test results.

| HrHPV test |  | Aptima |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Result | Positive |  | Negative |  | In total |  |
|  |  | N | % | N | % | N | % |
| HC2 | Positive | 187 | 39.0 | 15 | 3.1 | 202 | 42.1 |
|  | Negative | 19 | 3.9 | 259 | 54.0 | 278 | 57.9 |
|  | In total | 206 | 42.9 | 274 | 57.1 | 480 | 100.0 |

Table 2. HC2 results and histological status of the patients at the HPV sampling visit (N = 330 samples with histology). Performance parameters are shown for CIN2+.

|  | HC2+ |  | HC2- |  | Sensitivity | Specificity | PPV | NPV |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | N | % | N | % | 92.5% | 58.2% | 36.0% | 96.8% |
| CIN1- | 110 | 33.3 | 153 | 46.4 |  |  |  |  |
| CIN2+ | 62 | 18.8 | 5 | 1.5 |  |  |  |  |

CIN = cervical intraepithelial neoplasia

CIN1- = Normal and CIN1 findings

CIN2+ = CIN2 and higher grade lesions, prevalence 20.3% (N = 67).

Table 3. Aptima results and histological status of the patients at the HPV sampling visit (N = 330 samples with histology). Performance parameters are shown for CIN2+.

|  | Aptima+ |  | Aptima- |  | Sensitivity | Specificity | PPV | NPV |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | N | % | N | % | 94.0% | 59.3% | 37.1% | 97.5% |
| CIN1- | 107 | 32.4 | 156 | 47.3 |  |  |  |  |
| CIN2+ | 63 | 19.1 | 4 | 1.2 |  |  |  |  |

CIN = cervical intraepithelial neoplasia

CIN1- = Normal and CIN1 findings

CIN2+ = CIN2 and higher grade lesions, prevalence 20.3% (N = 67).

Table 4. Sensitivity, specificity and positive and negative predictive values of the hrHPV assays for the detection of CIN2+ stratified according to the reason for colposcopy.

| HrHPV assay | Reason for colposcopy | Sensitivity (%) | Specificity (%) | PPV (%) | NPV (%) |
| --- | --- | --- | --- | --- | --- |
| HC2 | 1 | 94.3 | 49.3 | 41.7 | 95.8 |
|  | 2 | 83.3 | 69.8 | 23.8 | 97.4 |
|  | 3 | 100.0 | 57.9 | 20.0 | 100.0 |
| Aptima | 1 | 94.3 | 51.4 | 42.7 | 95.9 |
|  | 2 | 91.7 | 69.8 | 25.6 | 98.7 |
|  | 3 | 100.0 | 57.9 | 20.0 | 100.0 |

1 = Patients who had Pap test suggestive of cervical lesion, N = 191

2 = Patients who had a treatment or control visit at which histology was obtained, N = 118

3 = Other patients, N = 21

Table 5. HrHPV results and cytological status of the patients at the HPV sampling visit (N = 480).

|  |  |  | HC2 |  |  |  | Aptima |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Positive |  | Negative |  | Positive |  | Negative |  |
| Cytology | N | %a | N | % | N | % | N | % | N | % |
| Normal | 261 | 56.4 | 54 | 20.7 | 207 | 79.3 | 59 | 22.6 | 202 | 77.4 |
| ASC-US | 77 | 16.6 | 31 | 40.3 | 46 | 59.7 | 32 | 41.6 | 45 | 58.4 |
| LSIL | 65 | 14.0 | 53 | 81.5 | 12 | 18.5 | 50 | 76.9 | 15 | 23.1 |
| ASC-H+b | 60 | 13.0 | 57 | 95.0 | 3 | 5.0 | 57 | 95.0 | 3 | 5.0 |
| No datac | 17 |  | 7 | 41.2 | 10 | 58.8 | 8 | 47.1 | 9 | 52.9 |
| Total | 480 |  | 202 | 42.1 | 278 | 57.9 | 206 | 42.9 | 274 | 57.1 |

a Proportion out of 463 samples with cytology

b ASC-H and more severe findings, including HSIL

c No cytological sample was taken on the HPV sampling visit

ASC-H = atypical squamous cells, cannot exclude HSIL

ASC-US = atypical squamous cells of undetermined significance

HSIL = high-grade squamous intraepithelial lesion

LSIL = low-grade squamous intraepithelial lesion

Table 6. Sensitivity, specificity and positive and negative predictive values of the hrHPV assays for the detection of CIN2+ in different cytological groups.

| HrHPV assay | Cytology | Sensitivity (%) | Specificity (%) | PPV (%) | NPV (%) |
| --- | --- | --- | --- | --- | --- |
| HC2 | LSIL- | 81.5 | 61.6 | 19.5 | 96.7 |
|  | ASC-H+ | 100.0 | 15.8 | 70.4 | 100.0 |
| Aptima | LSIL- | 85.2 | 62.9 | 20.7 | 97.4 |
|  | ASC-H+ | 100.0 | 15.8 | 70.4 | 100.0 |

ASC-H+ = ASC-H and more severe findings, including HSIL (N = 57)

LSIL- = Normal, ASC-US and LSIL findings (N = 264)