Continuous non-invasive measurement of cardiac output in neonatal intensive care using regional impedance cardiography: a prospective observational study

Banerjee J (MD (Res); FRCPCH)1,2,3,4, Khatib N2,3 (PhD), Mansfield RC (BA (Hons); BM BCh)1,2,3, Sathiyamurthy S (MD, FRCPCH)1, Kariholu U (MD, FRCPCH)1, Lees C (MD; FRCOG)2,3,5

Affiliations:

- Department of Neonatology, Queen Charlotte’s and Chelsea Hospital, Imperial College Healthcare NHS Trust, London, UK

- Biomedical Research Centre, Imperial College Healthcare NHS Trust

- Institute of Reproductive and Developmental Biology, Imperial College London, London, UK

- Origins of Health and Disease, Centre for Child Health, Imperial College London, London, UK

- Department of Fetal Medicine, Queen Charlotte’s and Chelsea Hospital, Imperial College Healthcare Trust, Du Cane Rd, White City, London W12 0HS

Short title: Non-invasive cardiac output in neonatal intensive care

Address correspondence and reprint requests to:

Dr Jayanta Banerjee, Department of Neonatology, Room 524 Hammersmith House, Queen Charlotte’s and Chelsea Hospital, Du Cane Road, London, United Kingdom, W12 0HS

Email: j.banerjee@imperial.ac.uk

Telephone: 020 3313 7308

Word count (excluding Title Page, Abstract, References and Figure Legend): 2500/2500

What is already known on this topic

The first three to four days after birth, often called the transitional period, are marked by significant haemodynamic changes. Current continuous haemodynamic monitoring in neonatal intensive care is limited to heart rate, arterial oxygen saturation, and blood pressure measurements. Bedside echocardiography, the only practical cardiac output (CO) assessment technique in neonatology, is performed intermittently and requires significant expertise. There are no validated continuous CO assessment technique and there are no widely accepted normative cardiac output values in neonates.

What this study adds

This prospective, observational study demonstrates the validity and feasibility of regional impedance cardiography (RIC) for monitoring CO continuously in preterm- and term-born neonates in an intensive care setting. Indicative normative measures of CO constructed in this study could be applicable to routine neonatal care. Neonates of lower gestational ages appear to have higher weight-adjusted cardiac output over the immediate and later postnatal period, when compared to those who were born at later gestational ages and those born at full term.

How this study might affect research, practice, or policy

This study provides clinicians and researchers with promising evidence for a technology that has the potential to transform haemodynamic assessment in the neonatal unit by allowing direct CO monitoring rather than relying on indirect measures.  These indicative normative measurement charts should be tested in future studies and clinical trials to see if continuous CO measurement and informed CO-guided haemodynamic management can reduce brain haemorrhage, or mortality in preterm neonates requiring intensive care.

Abstract (248/250 words)

Objectives

To compare agreement between echocardiography and regional impedance cardiography (RIC)-derived cardiac output (CO). To construct indicative normative ranges of CO for gestational age groups.

Design, Setting and Participants

Prospective cohort observational study performed in a tertiary centre in London, UK, including neonates born between 25-42 weeks gestational age.

Exposures

Neonates on the postnatal ward had two hours of RIC monitoring; neonates in intensive care had RIC monitoring for the first 72 hours, then weekly for two hours, with concomitant echocardiography measures.

Main outcomes and measures

RIC was used to measure CO continuously. Statistical analyses were performed using R (v4.2.2; R Core Team 2022). RIC-derived CO and echocardiography-derived CO were compared using Pearson’s correlations and Bland-Altman analyses. Differences in RIC-derived CO between infants born extremely, very, and late preterm were assessed using ANOVAs and mixed effects modelling.

Results

127 neonates (22 extremely, 46 very, 29 late preterm and 30 term) were included. RIC and echocardiography-measured weight-adjusted CO were correlated (r=0.62, p<0.001) with a Bland-Altman bias of -31ml/min/kg (limits of agreement -322 to 261ml/min/kg). The RIC-derived CO fell over 12 hours, then increased until 72 hours after birth. The 72-hour weighted-adjusted mean CO was higher in extremely preterm (424±158ml/min/kg) compared to very (325±131ml/min/kg, p<0.001), and late preterm (237±81ml/min/kg, p<0.001) neonates; this difference disappeared by 2-3 weeks of age.

Conclusions

RIC is valid for continuous, non-invasive CO measurement in neonates. Indicative normative CO ranges could help clinicians to make more informed haemodynamic management decisions, which should be explored in future studies.

Trial registration details

The study was registered with ClinicalTrials.gov (NCT04064177).

Introduction

Cardiovascular instability during the transitional period after birth makes neonates susceptible to brain and pulmonary hemorrhages, resulting in serious acute and long-term health consequences(1–3). Immature cerebral autoregulation, fluctuations in blood pressure (BP), and use of inotropic agents are associated with development of brain hemorrhages(4). To guide hemodynamic management in this period, neonatologists are heavily reliant on heart rate, peripheral arterial oxygen saturation, BP, and serum lactate and urine output. A fundamental determinant of perfusion and BP is cardiac output (CO). Echocardiography can be used to guide use of fluid boluses, inotropes, and vasopressors(5). However, echocardiography is performed intermittently, and suitable expertise is not always readily available(6). Gold-standard methods to measure CO (such as thermodilution(7), ultrasound dilution(8) and dynamic MRI(9)) are not clinically practical.

A validated, continuous, non-invasive CO monitoring technique could transform neonatal care, allowing clinicians to consider trends in CO alongside other measures before commencing vasoactive or fluid therapies. A systematic review and meta-analysis found promising results when comparing bioimpedence-based devices with routine CO measurement techniques in children and adolescents (n>1200), but not in neonates (n>700).(10) Most devices tested thus far in neonates and children use multiple sensor pads placed across the upper body (transthoracic electrical bioimpedance), rendering them impractical in very (<1500 grams) and extremely (<1000 grams) low birth weight neonates(10).

In this study, we assessed the validity of continuous regional impedance cardiography (RIC)-derived CO measurement using two sensor pads placed on a wrist and ankle, compared to echocardiography-derived CO measurement in neonates. Additionally, we created indicative normative longitudinal measures of CO in preterm and term neonates.

Materials and methods

The Non-invasive Continuous Cardiac Output Monitoring (NICCOM) study was conducted in a tertiary perinatal centre in London, UK. The study received approval from London-Bromley Research Ethics committee (19/LO/1290) and the protocol was pre-registered (https://clinicaltrials.gov/ct2/show/NCT04064177). The study aimed to assess validity of RIC for continuous CO monitoring compared to echocardiography-derived CO in neonates. RIC utilises whole body electrical bioimpedance rather that thoracic bioimpedance used in other neonatal studies. Secondly, the study aimed to create normative measures of cardiac output in neonates born between 25 weeks to 42 weeks of gestational age (GA) over the first 72 hours of life. For this second workstream, neonates were grouped by gestational age: extremely preterm neonates (those born up to 27+6 weeks GA); very preterm (28+0 to 31+6 weeks GA); late preterm (32+0 to 36+6 weeks GA); and term (>37+0 weeks GA)(11).

All stable preterm- and term-born neonates on the postnatal ward or admitted to the neonatal unit were eligible. Neonates who were very unstable and moribund and/or more than 12 hours old at the time of screening for eligibility, those with life-limiting congenital conditions or where the skin was very fragile were excluded. Prospective parental consent was obtained for neonates on the postnatal ward and deferred parental consent was obtained for those admitted to the neonatal unit. For neonates on the postnatal ward, two hours of RIC monitoring accompanied by echocardiography within the same period was performed daily until discharge. For the neonates on the neonatal unit, 72 hours of RIC monitoring was commenced within 12 hours of birth, and one echocardiography was performed every 24 hours in this period. Admitted neonates subsequently had weekly RIC monitoring for two hours, accompanied by echocardiography, until discharge or 40 weeks of corrected GA.

Cardiac output measurement

Standard echocardiographic technique(6) using GE Vivid S5/S70/E95 ultrasound machines was used to measure left ventricular SV and CO; performed by three neonatal consultants (JB, SS and UK) who are trained in performing neonatal echocardiography. All CO measurements were calculated offline by single investigator (JB). An RIC device called the non-invasive cardiac system (NICaS, NI Medical, Petah Tikvah, Israel) was used for continuous stroke volume (SV) and CO measurements. Left ventricular CO was selected amongst the other measures of systemic perfusion measurements such as right ventricular output and superior vena cava flow; RIC-derived CO has been validated with thermodilution measurements(12). Birth weight was used in the first 72 hours; for subsequent weekly measurements, the most recent weekly weight was used for weight adjustment of CO (ml/min/kg) and SV (ml/kg).

The NICaS RIC device utilises two paired impedance electrodes, placed on two limbs (usually a wrist and contra-lateral ankle), and operates on an alternating current of 1.4mA and 32kHZ (Supplementary figure 1). The technology exploits the principle that the electrical resistance of the blood is lower than that of surrounding tissues. Consequently, with each arterial systolic expansion (pulsation), a reduction in electrical resistance occurs. The analog resistance signals are received by the device, where they are amplified and filtered, then digitised and analysed via proprietary algorithms.(13)

Sample size

For generating the longitudinal chart of indicative normative ranges of CO, we planned to recruit 108 neonates (6 for each gestational week between 25 and 42 weeks); in addition, based on annual admissions a pragmatic sample of 40 preterm neonates with fetal growth restriction (FGR) was also planned. The sample size per gestational week and methodology are based on those of previously published standard fetal Doppler(14) and estimated fetal weight(15) reference charts. In both, an average of 5 observations per gestational week were used for the construction and evaluation of normal ranges. Allowing for loss of data in 20% cases, we chose 6 neonates per gestational week in this study.

Statistics

Analyses were executed using R Statistical Software (v4.2.2; R Core Team 2022). NICaS and echocardiography measurements were compared using Pearson’s product-moment correlation and visualised using correlation plots (Plotly for R, v4.10.0; Sievert, 2020). Bland-Altman plots were produced (ggplot2 package for R, v3.3.0) to assess measurement error and relative percentage error between NICaS and echocardiography. Equivalence between NICaS and echocardiography measurements was assumed if the mean percentage error (MPE) was <30% which is an accepted norm in most device studies.

Longitudinal NICaS data was displayed both continuously (6h - 72h post-birth) and in 6 hour and 7-day intervals, using means and standard deviations and standard errors (SE). A cubic spline was applied for smoothing continuous data (Stats package for R, v3.6.2). One-way ANOVA with Tukey’s post-hoc tests was employed to cross-sectionally examine statistical differences in NICaS measurements between gestational age groups.

Differences in weight-normalised CO progression and its comparison across gestational age groups was evaluated using mixed effects modelling (lme4 package for R, v2.3.2). Linear and quadratic curves were sequentially added to a random intercepts model and model fit assessed using log likelihood. The mixed effects model considered a fixed group effect and group * time interactions to examine the impact of gestational age group on CO changes over time, sequentially adding fixed time effects and testing fit at each stage.

Results

Patient characteristics

Neonatal and maternal characteristics and core neonatal outcomes are detailed in Table 1 and 2, respectively. Three neonates died, and two had severe (Grade 3) intraventricular haemorrhage (IVH); these neonates were excluded from the normative range analysis. 127 neonates were included in the normative range analysis: 22 extremely preterm (25+0 to 27+6 weeks  GA); 46 very preterm (28+0-32+6) weeks GA; 29 late preterm (33+0-36+6) and 30 term (>37+0). Of the 367 paired RIC and echocardiography measurements, 34 were obtained during invasive ventilation; 189 during non-invasive ventilation; and 144 obtained from neonates not on respiratory support.

Comparison between Echocardiography- and RIC-derived CO

Echo- and RIC-derived weight-normalised CO measurements showed a moderate positive correlation (r=0.62, p<0.001) (Fig 1). For all measurements, the bias was -31ml/min/kg with LoA from -322 to +261ml/min/kg and a mean percentage error (MPE) of 4.1% (Fig 2). Analysis of the bias and LoA across independent gestational age groups revealed that LoAs became narrower, as gestational age increased (Supplementary figure 2).

Comparison of CO over time between gestational age groups

The median and IQR of continuous CO measurements showed distinctive patterns over the first 72 hours in the gestational age groups (Supplementary figure 3). The continuous CO and SV levels were overall higher in neonates of lower gestational ages (Fig 3). The 72-hour weighted-adjusted mean CO was higher in extremely preterm (424±158ml/min/kg) compared to very preterm (325±131ml/min/kg, p<0.001), and late preterm (237±81ml/min/kg, p<0.001) neonates. There was an initial fall in CO and SV in the first 12h in all gestational age groups (Fig 3). Between 24h and 60h post-birth, CO and SV steadily increased in the extremely preterm neonates, while remaining stable for very and late preterm neonates. The initial fall (0-24h) and final peak of CO (60-72h) coincided with a similar SV and HR fluctuation (Supplementary figure 4).  The mean CO, SV and HR were higher (p<0.05) in earlier gestational age groups at 6, 12, 24, 48 and 72h post-birth, except for HR between extremely preterm and very preterm neonates (Supplementary figure 4). CO and SV were significantly higher in extremely preterm neonates compared to later gestational age babies at 7d and 14d post-birth (Supplementary figure 4); this correlated with the PDA sizes in these groups, with significant difference noted on day 3 (Kruskal-Wallis χ²(2)=13.102, p=0.001) and 7 (χ²(2)=13.207, p=0.001); and then no difference by 14 days (χ²(2)=1.396, p=0.498).

Using a mixed effects model, the intercept for extremely preterm neonates was higher than very preterm neonates but not statistically significant (Supplementary figure 5 and document 6), whereas the intercept for extremely preterm neonates was significantly higher than that of late preterm neonates (p<0.001). The increment in CO after 12 hours post birth was steeper in extremely preterm neonates (1.8 ml/min/kg per hour) compared to very (0.7 ml/min/kg per hour, p<0.05) and late preterm neonates (0.3 ml/min/kg per hour, p<0.01).

Discussion

This study demonstrated that monitoring CO continuously using RIC in neonates  is feasible and valid when compared with echocardiography. The probes were placed by trained team members. No skin reaction or damage occurred. We found that extremely preterm neonates have higher RIC-derived CO over the first 72 hours after birth and in the later postnatal weeks, when compared to term and very or late preterm babies. Gestational age and birthweight explained a large proportion of variance in CO as a function of time, and both showed an inverse correlation with CO.  CO fell in the immediate postnatal period, followed by a steady increase. Lower gestational age neonates showed a steeper rate of CO increase after 12 hours, followed by a shallower deceleration, compared to higher gestational age neonates. While late preterm neonates had similar weight-adjusted CO at 12 hours and 72 hours after birth, extremely preterm neonates had a 35% higher CO at 72 hours compared to 12 hours, indicating that they sustained higher CO levels for longer after birth.

Electrical bioimpedance has been validated for CO measurement in paediatric patients but not in neonates(10). Our meta-analysis comparing echocardiography with transthoracic bioimpedance(10) showed that the median (range) bias (L/min) in five neonatal studies was –0.01 L/min (–0.21 to 0.006 L/min), with median percentage error (PE) 45% (29–53%). When normalised by weight, the median bias (mL/min/kg) in four neonatal studies was 3.4 mL/min/kg (–8.9 to 18.8 mL/min/kg), and median PE 37% (24–60%). In the current study, using total body or regional impedance cardiography (RIC),   the bias was -31ml/min/kg (-12.6%) and the mean PE was 4.1% thereby demonstrating low systematic bias between RIC and bedside echocardiography derived CO. However, the limits of agreement of -320 to +260 ml/min/kg demonstrate variability. A large amount of the captured variability is likely owed to physiological variants of PDA and PFO, motion artifacts, and the impact of movement on HR and therefore CO. While it is not possible to conclude which of the two methods for measuring CO measurement is “more true” from this study, the unique benefits of RIC in allowing continuous, non-invasive measurement in practice and its minimal systematic bias against the current “gold standard” should be considered in its acceptability for adoption into clinical practice. It is also notable that the limits of agreement became narrower with increasing gestational age, thereby indicating potentially greater equivalence between devices in infants closer to term age.

Murase et al. used intermittent bedside echocardiography to measure 12 hourly CO in 109 preterm neonates between 22-33 weeks of gestational age(16). A post-hoc analysis of CO of neonates born before 26 weeks of gestational age, compared to those who were born at a higher gestational age, showed a fall in CO at 3 and 12 hours after birth, similar to our findings.14 In our study, the decline was most prominent in extremely preterm neonates. Immediately after birth as the lung inflates, pulmonary vascular resistance (PVR) drops with resultant increase in stroke volume and cardiac output(17). Animal studies have demonstrated that birth is associated with sudden catecholamine surge increasing HR(18,19) and increase in thyroid hormone release(20), resulting in increased cardiac contractility(21). After 6-12 hours there is a drop in this surge and therefore decrease in heart rate, contractility, and systemic vascular resistance (SVR). We can speculate that this fall in CO noted in the current study could be attributable to this physiologic haemodynamic adaptation. SVR is reported to be higher in extremely preterm neonates and there is a well-recognised fall in BP within the first 12-24 hours in extremely and very preterm neonates(22). Large retrospective epidemiological studies show that both low BP and inotropic support increased the risk of developing brain injury and IVH(23). Randomised controlled trials (RCT) trying to increase BP using inotropes such as dopamine have failed to demonstrate a reduction of IVH or mortality in preterm neonates(24,25). We hypothesise based on the present study that the decline in SVR, CO and therefore BP in the first 6-12 hours of age may be physiologically normal in preterm neonates, therefore, judicious use of vasopressor or inotropic agents is warranted.

Limitations

Inter-operator variability of echocardiography can be >20%(26,27). We acknowledge this as a limitation of this study, but bedside echocardiography is the only currently available practical technique to measure cardiac function in neonates. Several factors may affect the agreement between echo and RIC, including the stimulation of the neonate during echo resulting increase the HR or cause an increased amount of motion artefact in the RIC signal. Additionally, RIC measures a peripheral impedance waveform, so not equivalent to echo, which measures CO at the outflow tract(12) Only stable neonates in a single tertiary neonatal unit were included to aim for ‘normative measures’ of CO, limiting the generalisability of our findings. Although the numbers are small, this is the first study demonstrating continuous longitudinal CO normative measures over the transitional period and weekly measures up to four weeks of postnatal age.

Conclusions

The current study demonstrates that RIC can be used to measure neonatal CO non-invasively and continuously. In conjunction with BP nomograms, the indicative normative measures could help clinicians to understand haemodynamic status in neonatal intensive care, which should be tested in future research. We propose that RIC could be used in future studies to interrogate the haemodynamic effects of routine interventions and vasopressors; and to explore the relationship of CO with development of IVH. This work could pave way to developing RCTs of haemodynamically guided therapy aiming to reduce or prevent morbidity and mortality in neonates requiring intensive care.

Contributors

JB and CCL conceived and designed this study, oversaw study setup, conduct, and analysis setup, contributed to the first draft, approved the final version, and take responsibility for the accuracy of reported findings. SS, RM and UK contributed to study design, setup, conduct, contributed to the first draft, and approved the final version. NK wrote the statistical analysis plan and performed all the data analyses, contributed to the first draft, and approved the final version. JB, SS, RM and UK recruited patients, collected data, contributed to manuscript writing, and approved the final version. All authors had full access to the presented data in this Article and had final responsibility for the decision to submit for publication.

Declaration of interests

None of the authors have any competing interests to declare and have not benefited financially or by other means from the study. The NICaS study equipment was provided by NI Medical for the duration of the study but they did not have any role in study design/data analysis/interpretation.

Data sharing

Deidentified participant data will be shared upon reasonable request. Requests need to be addressed to the corresponding author.

Acknowledgements

We want to thank the parents and the babies who participated to the NICCOM study. JB and CCL are supported by the UK National Institute for Health Research (NIHR) Biomedical Research Centre (BRC) based at Imperial College Healthcare National Health Service Trust and Imperial College London. RM was supported by the Imperial BRC and the Imperial Health Charity grant (RFPR2021_29). NK received funding from the Imperial BRC. This study was funded through the Imperial NIHR Biomedical Research Centre. This article is independent research funded by the NIHR BRC, and the views expressed in this publication are those of the authors and not necessarily those of the NHS, NIHR, or the Department of Health. None of the funders have had any influence over study design, collection, analysis, and interpretation of the data, in writing the report and in the decisions to submit this article for publication. Non-invasive cardiac system (NICaS) equipments was supplied by NI Medical, Petah Tikvah, Israel. We would like to thank Ms Batia Gourin, Ms Zoe McClure and Ms Kirupalani Mariampillai, the research nurses and all the neonatal departmental nurses for their active support in this study. The authors would like to thank Prof John Cleland for his support and guidance in writing this manuscript.

References

1.	Costeloe K, Hennessy E, Gibson AT, Marlow N, Wilkinson AR. The EPICure Study: Outcomes to Discharge From Hospital for Infants Born at the Threshold of Viability. Pediatrics [Internet]. 2000 Oct 1;106(4):659–71. Available from: https://publications.aap.org/pediatrics/article/106/4/659/65772/The-EPICure-Study-Outcomes-to-Discharge-From

2.	Costeloe KL, Hennessy EM, Haider S, Stacey F, Marlow N, Draper ES. Short term outcomes after extreme preterm birth in England: comparison of two birth cohorts in 1995 and 2006 (the EPICure studies). BMJ [Internet]. 2012 Dec 4;345(dec04 3):e7976–e7976. Available from: https://www.bmj.com/lookup/doi/10.1136/bmj.e7976

3.	Moore T, Hennessy E, Myles J, al. et. Neurological and developmental outcome in extremely preterm children born in england in 1995 and 2006: the EPICure studies. BMJ. 2012;345:e7961.

4.	Cimatti AG, Martini S, Galletti S, Vitali F, Aceti A, Frabboni G, et al. Cerebral Oxygenation and Autoregulation in Very Preterm Infants Developing IVH During the Transitional Period: A Pilot Study. Front Pediatr. 2020 Jul 15;8.

5.	Vrancken SL, van Heijst AF, de Boode WP. Neonatal Hemodynamics: From developmental physiology to comprehensive monitoring. Vol. 6, Frontiers in Pediatrics. Frontiers Media S.A.; 2018.

6.	de Boode WP, Kluckow M, McNamara PJ, Gupta S. Role of neonatologist-performed echocardiography in the assessment and management of patent ductus arteriosus physiology in the newborn. Seminars in Fetal and Neonatal Medicine. 2018;

7.	Székely A, Breuer T, Sápi E, Székely E, Szatmári A, Tóth M, et al. Transpulmonary Thermodilution in Neonates Undergoing Arterial Switch Surgery. Pediatr Cardiol [Internet]. 2011 Feb 9;32(2):125–30. Available from: http://link.springer.com/10.1007/s00246-010-9828-0

8.	Vrancken SL, de Boode WP, Hopman JC, Singh SK, Liem KD, van Heijst AF. Cardiac output measurement with transpulmonary ultrasound dilution is feasible in the presence of a left-to-right shunt: a validation study in lambs. Br J Anaesth [Internet]. 2012 Mar;108(3):409–16. Available from: https://linkinghub.elsevier.com/retrieve/pii/S0007091217323048

9.	Ficial B, Finnemore AE, Cox DJ, Broadhouse KM, Price AN, Durighel G, et al. Validation Study of the Accuracy of Echocardiographic Measurements of Systemic Blood Flow Volume in Newborn Infants. Journal of the American Society of Echocardiography [Internet]. 2013 Dec;26(12):1365–71. Available from: https://linkinghub.elsevier.com/retrieve/pii/S0894731713006615

10.	Mansfield RC, Kaza N, Charalambous A, Milne AC, Sathiyamurthy S, Banerjee J. Cardiac Output Measurement in Neonates and Children Using Noninvasive Electrical Bioimpedance Compared with Standard Methods: A Systematic Review and Meta-Analysis. Crit Care Med. 2022 Jan 1;50(1):126–37.

11.	Blencowe H, Cousens S, Oestergaard MZ, Chou D, Moller AB, Narwal R, et al. National, regional, and worldwide estimates of preterm birth rates in the year 2010 with time trends since 1990 for selected countries: a systematic analysis and implications. The Lancet. 2012 Jun;379(9832):2162–72.

12.	Cotter G, Schachner A, Sasson L, Dekel H, Moshkovitz Y. Impedance cardiography revisited. Physiol Meas. 2006 Sep 1;27(9):817–27.

13.	Cotter G, Schachner A, Sasson L, Dekel H, Moshkovitz Y. Impedance cardiography revisited. Physiol Meas. 2006 Sep 1;27(9):817–27.

14.	Parra-Cordero M, Lees C, Missfelder-Lobos H, Seed P, Harris C. Fetal arterial and venous Doppler pulsatility index and time averaged velocity ranges. Prenat Diagn. 2007 Dec 30;27(13):1251–7.

15.	Hadlock FP, Harrist RB, Sharman RS, Deter RL, Park SK. Estimation of fetal weight with the use of head, body, and femur measurements—A prospective study. Am J Obstet Gynecol. 1985 Feb;151(3):333–7.

16.	Murase M, Ishida A, Momota T. Serial Pulsed Doppler Assessment of Early Left Ventricular Output in Critically Ill Very Low-Birth-Weight Infants. Pediatr Cardiol [Internet]. 2002 Jul 24;23(4):442–8. Available from: http://link.springer.com/10.1007/s00246-002-1457-9

17.	Azhibekov T, Noori S, Soleymani S, Seri I. Transitional cardiovascular physiology and comprehensive hemodynamic monitoring in the neonate: Relevance to research and clinical care. Vol. 19, Seminars in Fetal and Neonatal Medicine. 2014.

18.	Mehandru PL, Assel BG, Nuamah IF, Fanaroff AA, Kalhan SC. Catecholamine Response at Birth in Preterm Newborns. Neonatology. 1993;64(2–3):82–8.

19.	Padbury J, Agata Y, Ludlow J, Ikegami M, Baylen B, Humme J. Effect of fetal adrenalectomy on catecholamine release and physiologic adaptation at birth in sheep. Journal of Clinical Investigation. 1987 Oct 1;80(4):1096–103.

20.	Sack J, Beaudry M, Delamater P V, Oh W, Fisher DA. Umbilical Cord Cutting Triggers Hypertriiodothyroninemia and Nonshivering Thermogenesis in the Newborn Lamb. Pediatr Res. 1976 Mar;10(3):169–75.

21.	Riemenschneider TA, Brenner RA, Mason DT. Maturational Changes in Myocardial Contractile State of Newborn Lambs. Pediatr Res. 1981 Apr;15(4):349–56.

22.	Batton B. Neonatal Blood Pressure Standards. Clin Perinatol. 2020 Sep;47(3):469–85.

23.	Faust K, Härtel C, Preuß M, Rabe H, Roll C, Emeis M, et al. Short-term outcome of very-low-birthweight infants with arterial hypotension in the first 24 h of life. Arch Dis Child Fetal Neonatal Ed. 2015 Jul 21;100(5):F388–92.

24.	Dempsey EM. Challenges in treating low blood pressure in preterm infants. Vol. 2, Children. MDPI; 2015. p. 272–88.

25.	Dempsey EM, Barrington KJ, Marlow N, O’Donnell CP, Miletin J, Naulaers G, et al. Management of hypotension in preterm infants (The HIP Trial): A randomised controlled trial of hypotension management in extremely low gestational age newborns. Neonatology. 2014;105(4):275–81.

26.	Mertens L, Seri I, Marek J, Arlettaz R, Barker P, McNamara P, et al. Targeted Neonatal Echocardiography in the Neonatal Intensive Care Unit: Practice Guidelines and Recommendations for Training. Journal of the American Society of Echocardiography. 2011 Oct;24(10):1057–78.

27.	Harada K, Shiota T, Takahashi Y, Tamura M, Toyono M, Takada G. Doppler Echocardiographic Evaluation of Left Ventricular Output and Left Ventricular Diastolic Filling Changes in the First Day of Life. Pediatr Res. 1994 Apr;35(4):506–9.

Table 1. Patient details of neonates recruited in each gestational age groups: extremely preterm neonates are those born upto 27+6 weeks of gestational age, very preterm 28+0 to 31+6 weeks, late preterm 32+0 to 36+6 weeks, and term >37+0 weeks of gestational age at birth.

|  | Extremely preterm | Very preterm | Late preterm | Term |
| --- | --- | --- | --- | --- |
| Gestational age (Median; IQR) | 26+3 (25+6-27+1) | 30+3 (28+7-31+0) | 34+1 (32+4- 35+6) | 39+1 (38+3-40+6) |
| Birth weight in kg (Median; IQR) | 0.78 (0.68-0.90) | 1.23 (1.06-1.42) | 1.78 (1.45-2.29) | 3.40 (2.96-3.74) |
| Sex (Female: Male) | 14:11 | 22:14 | 16:21 | 24:12 |
| White British:white non-British:non-white | 5:4:16 | 3:6:27 | 4:11:22 | 14:12:10 |
| Maternal age in years (Median; IQR) | 36 (32-41) | 33 (30-39) | 33 (31-37) | 35 (32-38) |
| Maternal PET/PIH (n; percentage) | 4 (16%) | 10 (28%) | 8 (22%) | 0 (0%) |
| Maternal chorioamnionitis (n; percentage) | 15 (60%) | 9 (25%) | 3 (8%) | 0 (0%) |
| Antepartum haemorrhage (n; percentage) | 3 (12%) | 2 (6%) | 2 (5%) | 0 (0%) |
| Antenatal steroids received (n; percentage) | 18 (72%) | 27 (75%) | 18 (49%) | 0 (0%) |
| Optimum cord clamping practised (n; percentage) | 5 (20%) | 17 (47%) | 12 (32%) | 14 (39%) |

PET – pre-eclampsia, PIH – pregnancy induced hypertension

Table 2. Patient details with core neonatal outcomes

|  | Extremely preterm | Very preterm | Late preterm | Term |
| --- | --- | --- | --- | --- |
| Duration of ventilation (days): Invasive (Median; IQR) | 3 (0.5-8.75) | 0 (0-0) | 0 (0-0) | 0 (0-0) |
| Non-invasive (Median; IQR) | 35 (8-49) | 5.5 (2-9) | 0 (0-2) | 0 (0-0) |
| Nasal cannula oxygen (Median; IQR) | 4 (0-18.75) | 0 (0-11.75) | 0 (0-0) | 0 (0-0) |
| All Intraventricular haemorrhage (n; percentage) | 15 (60%) | 9 (25%) | 1 (3%) | 0 (0%) |
| Intraventricular haemorrhage grade 3 (n; percentage) | 1 (4%) | 0 (0%) | 0 (0%) | 0 (0%) |
| Pulmonary haemorrhage (n; percentage) | 0 (0%) | 0 (0%) | 0 (0%) | 1 (3%) |
| Bronchopulmonary dysplasia (n; percentage) | 14 (56%) | 6 (17%) | 0 (0%) | 0 (0%) |
| Necrotising enterocolitis (n; percentage) | 2 (8%) | 0 (0%) | 0 (0%) | 0 (0%) |
| PDA requiring medical/surgical treatment (n; percentage) | 3 (12%) | 0 (0%) | 0 (0%) | 0 (0%) |
| Death (n; percentage) | 2 (8%) | 1 (4%) | 0 (0%) | 0 (0%) |

List of figures

Figure 1. Correlation between regional impedance cardiography (RIC) and echocardiography derived CO over the same time-period. Pearson’s correlation coefficient of 0.62 (p<0.001) demonstrates a moderate correlation between cardiac output measures derived from Regional Impedance Cardiography (RIC) and echocardiography, adjusted for body weight (ml/min/kg).

Figure 2. Bland-Altman analysis of agreement between RIC and echocardiography derived CO. Measurement bias of -31 ml/min/kg and limits of agreement of -322 to +261 ml/min/kg demonstrate minimal systematic bias, but with wide limits of agreement. Agreement between RIC and echocardiography CO decreased with increasing CO value.

Figure 3. Cross-sectional and longitudinal analyses of normative continuous measures derived from RIC, depicting CO, stroke volume, and heart rate across varying preterm gestational age groups. Infants in lower gestational age groups exhibited higher overall CO and stroke volume across the first 72 hours of postnatal age. All curves are expressed in mean (dark solid or dotted lines) ±1 standard deviation (shaded areas) of the infants for each gestational age group over the first 72 hours of postnatal age.

Supplementary materials

Supplementary figure 1. Bland-Altman analysis of agreement between RIC and echocardiography derived CO for each infant gestational age group. Increasing gestational age was associated with narrower limits of agreement.

Supplementary figure 2. Comparison of 6 hourly cardiac output, stroke volume and heart rate measured using RIC over the first 72 hours followed by weekly measures up to one month expressed in mean and standard errors of mean (SEM). The 6 hourly mean measurements up to 72 hours and subsequent mean weekly measurements are compared using continuous measures analysis of variance.

Supplementary figure 3. Change in weight-adjusted cardiac output between 12 - 72 hours. Mixed effects model with linear fit revealed increased cardiac output over time in neonates of younger gestational age.

Supplementary figure 4. Median and inter-quartile ranges of continuous CO for neonates in the gestational age groups using RIC from birth up to 72 hours of postnatal age: the three panels indicating the extremely preterm, very preterm and late preterm neonates.

Supplementary figure 5. Regional impedance cardiography (RIC) electrodes on a mannikin with the monitor demonstrating the various parameters such as stroke volume, heart rate, cardiac output. The arrows in the electrodes pointing towards the heart from both upper and lower extremity.