Screening with blood 7α-hydroxy-4-cholesten-3-one by liquid chromatography tandem mass spectrometry to identify SeHCAT-diagnosed bile acid diarrhoea.

James M Hawley, Jonathan S Atkins, Christopher Bannon, Evette Hillman, Brian Shine, Jack Satsangi, Julian RF Walters

Manchester University NHS Foundation Trust

Institute of Metabolic Science, Addenbrooke’s Hospital, Cambridge

Oxford University Hospitals,

Imperial College Healthcare London, UK

Introduction:

Bile acid diarrhoea (BAD) causes digestive symptoms including frequent loose stools, urgency and incontinence, but diagnosis is often delayed and patients remain untreated. The gold-standard investigation is the SeHCAT test, measuring bile acid retention and loss over 7 days. This has had availability problems and requires two visits to a nuclear medicine department.  A cheaper, screening test would increase diagnostic rates. 7α-hydroxy-4-cholesten-3-one (C4), measured in blood, quantifies new bile acid synthesis. A liquid chromatography tandem mass spectrometry (LC-MS/MS) assay has been developed and validated.  We aimed to investigate the performance of this assay to assess how this can help select patients for SeHCAT and reduce unnecessary tests.

Methods:

Patients (n=133, M:F= 40:93) from two separate centres (n=108, London and n=25, Oxford) referred for SeHCAT testing were recruited. All patients provided early-morning serum samples for C4 analysis. The diagnostic performance of serum C4 against SeHCAT was analysed, using a SeHCAT retention of ≤10%, conventionally representing severe or moderate BAD.

Results:

40 patients (30%) had a positive SeHCAT <10%, 19 (14%) mild 10-15%, 22 (17%) borderline 15-20%, and 52 (39%) were negative >20%. C4 levels were inversely proportional to SeHCAT retention. The median C4 was 103 nmol/L in patients with SeHCAT <10%, and 29 nmol/L in those >20%. A ROC curve analysis of C4 showed an AUC of 0.82 to detect a SeHCAT <10%.

Optimal positive and negative cut-off values were calculated. A C4 value >120 nmol/L had 96% specificity and 37% sensitivity, and C4 of 30 nmol/L was 35% specific and 97% sensitive.

Using these values, 53 (40%) patients would no longer need SeHCAT to make or rule out a possible diagnosis of moderate or severe BAD. 80 (60%) would still require a SeHCAT test. This would include patients with severe or persistent symptoms who may be false negatives, and those with mild or borderline BAD.

Conclusions:

A reliable C4 assay has been established and its use as a screening test validated. It will benefit the diagnosis of BAD, with high and low cut-offs predicting the SeHCAT result in around 40% of those currently undergoing testing. There will be significant cost-benefits from this strategy. Its ease of use will encourage screening and increase the number of patients diagnosed and treated for BAD.